Optimization of reverse transfection conditions for siRNA or sgRNA oligos

Authors: Laurent Ozbun/Gianluca Pegoraro

Date: 08/14/2021

Revision: 2

Overview

- Day -1:

  - SiRNA/sgRNA transfection controls are spotted and then air dried into a PerkinElmer Cell Carrier Ultra 384 well imaging plate using the ECHO525 acoustic liquid handler (D304).

  - The plate is then sealed, and stored at –30o C.

- Day 0:

  - Set up the reverse transfection.

  - Use the Multidrop dispenser (D304, cell culture hood) to add 20 ul of OPTIMEM

  - Use the ECHO525 to dispense RNAiMAX into wells already containing 20 ul of OPTIMEM.

  - Use Multidrop dispenser to add 20 ul of cells resuspended in  20 ul of OPTIMEM/20% FBS with cells.

  - Incubate at 37o C

- Day + 2 or 3:

  - Fix and Stain cells using the Bluewasher washer/dispenser (D304, chemical fume hood).

  - If not imaging right away, store plate at 4o C.

- Day X:

  - Warm plate to RT before imaging, otherwise condensation on bottom of plate will impede image acquisition.

  - Image plate.

  - Recover plate and store in YOUR fridge.  Discard at your discretion.

Reagents

- siRNA controls are diluted to a 5 uM concentration (Stock).

- sgRNA controls (Synthego sgRNA kit v2) are resuspended at 3 uM concentration (Stock) in water. A working solution at 0.3 uM in water is prepared fresh before spotting. Do not freeze and thaw siRNA or sgRNA oligos solutions more than 3-4 times.

- RNAiMAX or other transfection reagent.

Preparation

- Day1: Set up ECHO 525 dispense protocol for siRNA.

  - Run a simulation of ECHO 525 transfer to ensure proper well dispenses and total volume used.

- Day 2: Set up ECHO 525 dispense protocol for RNAiMAX.

  - Run a simulation of ECHO 525 transfer to ensure proper well dispenses and total volume used.

- Set up Multidrop dispenser.

  - Make up solutions

  - Prime cassette

- Set up BlueCat Bluewasher.

  - Make up solutions

  - Prime Bluewasher

Day 1:

- Add siRNA control (5 uM) or sgRNA controls (0.3 uM) to Labcyte source plate (LDV or PP, as appropriate)

  - Be aware of the Dead volume and Maximum Volume per well for the LDV or PP source plate.

- Centrifuge Labcyte source plate at RT for 2 minutes at 1200 RPM.

- Conduct a survey of siRNA volume in Labcyte source plate to ensure there is adequate volume.

- Spot 150 nl of siRNA 5 uM (0.8 pmoles per well, final conc. in 40 ul: 20 nM) or 325 nl of sgRNA 0.3 uM (0.08 pmoles per well, final conc. in 40 ul: 2 nM) per well of a 384- well imaging plate.

- Air dry siRNA in the laminar air flow hood (usually takes 30 to 150 minutes).

- Seal the plate with Alufoil and store at -30oC.

Day 2:

- Thaw imaging plate at room temperature.

- During this time, set up ECHO 525 for dispensing RNAiMAX from the Labcyte source plate with appropriate RNAiMAX volume.

- For ONE set of controls (No siRNA, Neg siRNA, Pos siRNA; or No sgRNA, Neg sgRNA, Pos sgRNA) you will need 6.3 ul total to dispense a matrix for 25, 50, 75, 100, 125 and 150 nl RNAiMAX.

  - See “RNAiMAX_Optimization_Template” Excel spreadsheet for layout.

- Centrifuge SOURCE PLATE at RT for 2 minutes at 1200 RPM.

- Centrifuge thawed IMAGE PLATE at RT for 2 minutes at 1200 RPM

- Run a simulation of ECHO525 transfer to ensure proper well dispenses and total volume used.

- Conduct a survey of RNiMAX volume in source plate to ensure there is adequate volume.  Remember to calculate the 4 ul dead volume/well into your final total volume in the source plate.

- Add 20 ul of OptiMEM to the IMAGE PLATE using the Multidrop.  Aseptic techniques apply.

- Start ECHO525 run.

  - Click BLUE arrow.

  - Click RUN on new window that appeared.

  - Click START on new window that appeared.

- Follow directions.

  - For source plate window, give it a name (SOURCE_PLATE) if desired.

  - Once Destination plate is requested, first place image plate onto tray.

  - Click the pointer in the BARCODE window so the barcode will register.

  - Scan in barcode of your imaging plate.  This action will cause the ECHO 525 to accept the destination plate.

- Wait for ECHO525 to finish dispensing.

- Follow window cues once finished.

- Seal source plate (NOT your image plate) with AluFoil and return to minus 30.

- Add the clear lid on your image plate.

- Incubate image plate at RT for 30 minutes in the laminar flow hood.

- During this incubation step, prep the cells for the multidrop dispenser.

- Dispense 20 ul of cells into imaging plate.

- Incubate at RT for 30 minutes.

- Place plate at 37oC and incubate for 72 hours.

Day 5:

- Fix the cells with PFA using the Bluewasher.

- Stain the cells with DAPI or continue with your endpoint visualization assay.

Day X:

- Acquire cell image with CV7000S or CV8000.