Seeding cells in the imaging plates using the Multidrop Combi

Authors: Laurent Ozbun and Gianluca Pegoraro

Date: 03/16/2021

Revision: 1

This protocol describes the procedure to dispense 40 ul of trypsinized cells suspension into the wells of an imaging plate.

Equipment used:

- ThermoFisher Multidrop Combi dispenser (Room 41/D304, cell culture hood)

- ThermoFisher Standard tube dispensing cassette (Cat# 24072671)

Reagents:

- 40 ml of 10% v/v bleach in a 50 ml conical tube.

- 2 X 40 ml of sterile water in 2 separate 50 ml conical tubes.

- 40 ml of 0.1% v/v Tween 20 in water in a 50 ml conical tube.

  - Add 40 ul Tween 20 to 40 ml water.

- 2 x 40 ml of sterile PBS in a 50 ml conical tube.

- Complete cell culture medium with 2X FBS.

- OPTIMEM serum-free medium.

Cells:

- At least 40 ml of cells resuspended in OPTI-MEM (Medium + 2X FBS) at the appropriate concentration in a 50 ml conical tube. For cell density testing:

  - Calculate 15 ml for initial priming, then 20 ml volume for an ENTIRE plate.

  - Add 40 ul per well.

- Count cells using an automated cell counting instrument. Always use the same instrument for downstream applications. The HiTIF is equipped with the BioRad TC20 Automated Cell Counter.

Tips and Tricks:

- Do NOT run the tubing dry between solution changes.  It is our experience that bubbles form at the nozzles and interferes with volume dispensing and sterility.

- Do NOT allow salted solutions (PBS, medium, etc) to “sit” in dispenser. Evaporation will result in salts clogging the small bored nozzles.

- Follow the solution sequence as specified with priming volumes.

- Beware of the BLACK cords.  If these wrap around the peristaltic pump wheel, you cannot attach the cassette.  Ensure these are outside of the peristaltic pump wheels. This is a very common problem.

- Submerge the end of the dispensing tubes to the bottom of the tube.

- Be mindful of where you grab the tubing as you aseptically sterilize your tubing.

Preliminary Steps:

- Prepare the solutions to sterilize the tubing.

- Use the keypad on the Multidrop to have the correct program displayed and ready to go.

Using the Multidrop in 4 parts:

PART 1:  Harvest cells from your flask(s).

- Centrifuge to pellet cells. (1,500 RPM for 5 minutes).

- Resuspend cells in 10 ml of complete medium.

- Conduct a cell count using the Auto Cell Counter.

- Resuspend cells at the appropriate cell density in 40 ml of complete medium.

- Final volume of 40 ml per Tube. The number of tubes is dependent on the total number of plates you are transfecting.

    - The first 40ml is for priming and for the first plate.

    - The next 40 ml are for 2 plates.  I do not try to squeeze a 3rd plate.

- Prepare an extra tube of cells, just in case.

  - Some indicative cell numbers:

    - 600 cells/40 ul per well		1.5x 104 cells/ml

    - 1,000 cells/40 ul per well		2.5 x 104 cells/ml

    - 1,500 cells/40 ul per well		3.75 x 104 cells/ml

    - 2,000 cells/40 ul per well		5 x 104 cells/ml

    - 2,500 cell/40 ul per well		6.25 x 104 cells/ml

Prepare the Multidrop by sterilizing the cassette tubing.

- Wear gloves.

- Mount the dispensing cassette on the MultiDrop.

- Remember to pull the blue shield all the way forward.

- Connect the drain tube to the aspirator and turn on the vacuum.

- Submerge the aspiration end of the cassette tubes in the bleach 10% tube, all the way to the bottom of the tube.

- Prime the cassette with 20 ml of bleach 10%. To prime, press the PRIME button on the keypad.

- Prime with 20 ml of water (Tube 1).

- IMPORTANT: change the water tube!

- Prime with 20 ml of water (Tube 2).

PART 3:  Dispense cell suspension to your plate.

- Prime the cassette with 20 ml of PBS.

- Add your plate to the stage.

- IMPORTANT: remove the plate lid!

- Vortex your cells to mix.

- Dispense 40 ul of cell suspension to each well. To dispense the 40 ul, press START on the keypad.

- Place the plate lid back onto the plate.

- Process the next plate as needed (i.e. repeat steps 11 to 15 included).

- Incubate plate(s) at RT for 30 minutes. (With lid on and in the hood).  Continue immediately to the next step.

PART 4: Decontaminate your cassette

- Prime the cassette with 20 ml of PBS.

- Prime the cassette with 20 ml of 0.1 Tween 20%.

- Prime the cassette with 20 ml of bleach 10%.

- Prime with 20 ml of water.

- IMPORTANT: change the water tube!

- Prime with 20 ml of water.

- Prime the cassette dry.

- Slide the blue shield back.

- Remove and place the cassette back in the box.

- Write down the estimated total volume used in the experiment (Including all washes)