# SOP_010: SBS Illumina Sequencing

Standard Operating Procedure for In Situ Sequencing-by-Synthesis | isabel.quintanilla@nih.gov, gianluca.pegoraro@nih.gov

Isabel Quintanilla

January 13, 2026

## 1 Purpose and Scope

### 1.1 Objective

To perform in situ sequencing using MiSeq Nano or NovaSeqX kits with 4- and 3-color chemistry for optical pooled screening (OPS) applications, enabling single-cell barcode/sgRNA identification through iterative incorporation, imaging, and cleavage cycles.

### 1.2 Scope

This protocol applies to prepared cell samples following library preparation and covers the complete sequencing workflow including primer hybridization, nucleotide incorporation, imaging, and cleavage for multiple cycles until all barcodes/sgRNAs are resolved.

### 1.3 Prerequisites

Completed conventional or T7 library preparation (SOP_009a or SOP_009b)

Training on high-temperature thermocycler operation

Knowledge of fluorescence microscopy and imaging systems

Understanding of sequencing chemistry principles

Experience with time-critical procedures

## 2 Materials and Equipment

### 2.1 Equipment Required

Flat-top thermocycler with accurate temperature control

Fluorescent microscope with appropriate filter sets

Precision pipettes and tips

Plate handling tools for hot surfaces

Timer with multiple alarms

Temperature monitoring device

### 2.2 Reagents and Consumables

| Reagent Name | Supplier | Catalog # | Storage | Final Concentration | Price ($) |
| --- | --- | --- | --- | --- | --- |
| Sequencing Primer (oSBS_CROPseq-v2-T7-Hybrid) | IDT | Custom | -20°C | 1 μM | 15.00 |
| SSC 20X RNase-free | ThermoFisher Scientific | AM9763 | RT | 2X | 85.99 |
| PBS-T (0.05% Tween-20) | Prepare fresh | - | 4°C | 0.05% Tween | - |
| MiSeq Reagent Nano Kit v2 (500 cycles) | Illumina | MS-103-1003 | -20°C | As supplied | 400.00 |
| NovaSeqX Reagent Kit | Illumina | Variable | -20°C | As supplied | Variable |
| Illumina PR2 Buffer | Illumina | Included with kit | RT | As supplied | Included |
| DAPI Staining Solution | ThermoFisher Scientific | D1306 | 4°C | 200 ng/mL | Variable |
| Isopropyl Alcohol (IPA) 99% | Various Suppliers | N/A | RT | 99% | Variable |
| Lens Cleaning Tissue | Various Suppliers | N/A | RT | - | Variable |

Note: You either use the MiSeq Reagent Nano Kit v2, or the NovaSeqX Reagent Kit, not both of them.

Note: Refer to OPS_Reagents_Master_List.csv and OPS_Oligonucleotides.csv for complete specifications.

Note: The NovaSeqX Reagent Kit reagent cannot be ordered from the vendor to be used in OPS. The Kit needs to be opened and used in a NovaSeqX DNA sequencer so that it is opened and mixed. After use in the instrument, collect the leftover reagents in the used kit and freeze at -80C until aliquoting in 50 ml volumes (The volume of the aliquots depends on the number of OPS cycles and the plate format used). ⚠️ Critical Point: Avoid repeated freeze/thaw cycles.

### 2.3 Software Requirements

Microscope control software with automated imaging capabilities

Image analysis software for sequencing data processing

Temperature monitoring software

## 3 Reagent Preparation

### 3.1 Sequencing Primer Solution

Final Volume: 10 mL Storage: Prepare fresh or use frozen aliquots Shelf Life: 6 months at -20°C

| Component | Stock Conc. | Final Conc. | Volume | Notes |
| --- | --- | --- | --- | --- |
| Sequencing primer | 100 μM | 1 μM | 100 μL | Custom designed |
| 2X SSC buffer | 2X | 2X | 9.9 mL |  |

### 3.2 DAPI Working Solution

Final Volume: 50 mL Storage: Prepare fresh Shelf Life: Same day use only

| Component | Stock Conc. | Final Conc. | Volume | Notes |
| --- | --- | --- | --- | --- |
| DAPI | 5 mg/mL | 200 ng/mL | 2 μL | Light sensitive |
| 2X SSC buffer | 2X | 2X | 49.998 mL |  |

## 4 Procedure

### 4.1 Primer Hybridization

Time Required: 45 minutes Temperature: 37°C

Primer Application Apply 1 μM sequencing primer in 2X SSC buffer to cover all wells ✅ Quality Check: Ensure complete coverage of sample area

Hybridization Incubate at 37°C for 30 minutes ⚠️ Critical Point: Maintain accurate temperature for proper hybridization

Washing Rapidly wash three times with PBS-T Remove all excess primer solution

### 4.2 Nucleotide Incorporation Cycle

Time Required: 15 minutes per cycle Temperature: 60°C ⚠️ CRITICAL TIMING: Do not exceed 5-minute incorporation limit

Incorporation Mix Application Pre-warm thermocycler to 60°C Apply incorporation mix to samples: MiSeq Nano kit: Use reagent 1 as supplied NovaSeqX kit: Use reagent 30 as supplied ⚠️ Critical Point: Work quickly to minimize temperature variation

Incorporation Incubation Place plate on flat-top thermocycler immediately Incubate for exactly 5 minutes at 60°C ⚠️ CRITICAL: Do not exceed 5 minutes - extending time greatly increases background

Rapid Dilution and Washing Remove plate from thermocycler Immediately add Illumina PR2 buffer to dilute incorporation mix Aspirate diluted mix quickly Add and remove PR2 buffer 6 times at room temperature ⚠️ Critical Point: Speed is essential - delayed washing increases background

### 4.3 High-Temperature Washing

Time Required: 30 minutes Temperature: 60°C

Intensive Washing Protocol Wash 5 times with PR2 buffer at 60°C Each wash: 5 minutes duration 🔧 Troubleshooting: If background issues persist, increase wash number and duration in later cycles

### 4.4 DAPI Staining and Imaging

Time Required: 15 minutes Temperature: Room temperature

Nuclear Staining Exchange into 200 ng/mL DAPI in 2X SSC Allow 5 minutes for DAPI binding to DNA ⚠️ Critical Point: Protect from light during incubation

Plate Preparation for Imaging Clean bottom of plate with isopropyl alcohol and lens cleaning tissue Visually inspect to ensure no streak marks remain ✅ Quality Check: Clear optics are essential for accurate imaging

Image Acquisition Use pre-set optical configuration specific to sequencing kit: MiSeq Nano: 4-color configuration NovaSeqX: 3-color configuration Acquire all necessary channels for the cycle ✅ Quality Check: Verify signal quality and minimal background Note Refer to SOP_011: Imaging Using Nikon Ti2 for additional details on how to setup the microscope.

### 4.5 Cleavage and Cycle Completion

Time Required: 15 minutes Temperature: 60°C

Cleavage Reaction Apply cleavage mix after imaging: MiSeq Nano kit: Use reagent 4 as supplied NovaSeqX kit v2: Use reagent 6 as supplied Incubate at 60°C for 6 - 12 minutes.

Post-Cleavage Washing Quickly wash 3 times with PR2 at room temperature Wash 3 times with PR2 at 60°C for 1 minute each Quickly wash 3 times with PR2 at room temperature ⚠️ Critical Point: Temperature transitions must be rapid ✅ Quality Check: Verify that all the signal is gone after washing

### 4.6 Cycle Repetition

Continue Sequencing Repeat steps 4-12 for subsequent cycles Continue until sufficient cycles completed to resolve all barcodes/sgRNAs 🔧 Planning: Determine cycle number based on library complexity

## 5 Quality Control and Validation

### 5.1 Critical Control Points

CCP 1: Primer hybridization temperature accuracy (37°C ± 1°C)

CCP 2: Incorporation timing exactly 5 minutes maximum

CCP 3: Immediate dilution and washing after incorporation

CCP 4: High-temperature wash effectiveness (60°C ± 1°C)

CCP 5: Clean plate optics before each imaging session

CCP 6: Cleavage timing accuracy (6 minutes at 60°C)

### 5.2 Expected Results

Clear, bright signals at expected locations

Minimal background fluorescence

Signal intensity consistent across cycles

Successful signal removal after cleavage

Preserved cell morphology throughout cycles

### 5.3 Cycle-Specific Monitoring

Cycle 1: Establish baseline signal intensity

Mid-cycles: Monitor background accumulation

Mid-cycles: Verify that all the signal is gone after washing

Final cycles: Confirm complete barcode resolution

## 6 Troubleshooting

| Problem | Possible Cause | Solution | Prevention |
| --- | --- | --- | --- |
| High background fluorescence | Extended incorporation time | Strictly limit to 5 minutes, increase wash intensity | Use timer alarms, multiple operators for speed |
| Poor signal intensity | Incomplete incorporation | Check reagent activity, verify temperature | Fresh reagents, calibrate thermocycler |
| Signal persistence after cleavage | Incomplete cleavage | Extend cleavage time, check reagent | Verify cleavage reagent activity |
| Blurry images | Plate surface contamination | Clean with IPA and lint-free tissue | Handle plates carefully, check before each cycle |
| Temperature fluctuations | Thermocycler malfunction | Monitor with independent thermometer | Regular equipment calibration |
| Increasing background in later cycles | Accumulating wash artifacts | Increase wash stringency progressively | Start with higher wash standards |

## 7 References and Related Documents

### 7.1 Related SOPs

SOP_009a: Conventional Library Preparation

SOP_009b: T7 Library Preparation

SOP_011: Imaging Using Nikon Ti2

### 7.2 Kit Documentation

MiSeq Nano Kit v2 User Guide (Illumina)

NovaSeqX Kit User Guide (Illumina)

Sequencing by Synthesis Chemistry Guide (Illumina)

### 7.3 Literature References

[In situ sequencing methodology reference]

[SBS chemistry optimization reference]

[Background reduction strategies reference]

## 8 Appendices

### 8.1 Timing Guidelines

Critical Time Limits:

Incorporation: Maximum 5 minutes

Post-incorporation dilution: <30 seconds

Temperature transitions: <2 minutes

Cleavage: 6 - 12 minutes, cell line dependent

### 8.2 Kit-Specific Parameters

MiSeq Nano (4-color):

Reagent 1: Incorporation mix

Reagent 4: Cleavage mix

Imaging: 4 fluorescent channels

NovaSeqX (3-color):

Reagent 30: Incorporation mix

Reagent 6: Cleavage mix

Imaging: 3 fluorescent channels

### 8.3 Imaging Configuration

Objective: High NA for maximum light collection

Exposure: Optimize for signal/background ratio

Focus: Maintain throughout heating/cooling cycles

Filters: Kit-specific wavelength sets