# SOP_011: Imaging Using Nikon Ti2 Microscope

Standard Operating Procedure for Phenotype and ISS Image Acquisition | isabel.quintanilla@nih.gov, gianluca.pegoraro@nih.gov

Isabel Quintanilla

January 13, 2026

## 1 Purpose and Scope

### 1.1 Objective

To perform high-quality phenotype and in situ sequencing (ISS) imaging using the Nikon Ti2 microscope system for Optical Pooled Screening (OPS) applications.

### 1.2 Scope

This protocol covers microscope startup, phenotype acquisition, ISS imaging setup, and shutdown procedures. It includes both phenotype characterization and T7-ISS imaging workflows.

### 1.3 Prerequisites

Basic microscopy training

Familiarity with NIS Elements software

Knowledge of optical configurations

Understanding of JOBS pipeline workflows

## 2 Materials and Equipment

### 2.1 Equipment Required

Nikon Ti2 microscope system

NIS Elements software

Objectives: 4X, 10X, 40X, 60X (water immersion)

Excitation/emission filter sets

Incubator chamber

Sample plate holder

Water immersion delivery system (WID)

### 2.2 Software Requirements

NIS Elements AR: Primary acquisition software

JOBS Explorer: Automated imaging pipeline management

### 2.3 Optical Configurations

| Application | Objective | Excitation | Emission | Exposure | Notes |
| --- | --- | --- | --- | --- | --- |
| DAPI | 10X/40X/60X | 360-370nm | 435-485nm | Variable | Used for segmentation and alignment |
| Phenotype GFP | 40X/60X | 460-490nm | 510-560nm | Variable | Water objective |
| MiSeq T7-ISS | 10X | Kit specific | Kit specific | Variable | Signal optimization |
| NovaSeqX T7-ISS | 10X | Kit specific | Kit specific | Variable | Signal optimization |

## 3 System Startup Procedure

### 3.1 Power-Up Sequence

Time Required: 15 minutes Temperature: Room temperature

Main Power Switch Turn main power switch ON (located behind microscope monitor with orange sticker) ✅ Check: Power indicator LED illuminated

Peripheral Systems Turn incubator ON Turn lamp boxes ON ✅ Check: All status lights green

Microscope Systems Turn main microscope switch ON (under microscope monitor) Turn microscope and digital mirror switch ON ✅ Check: System initialization complete

Software Initialization Start NIS Elements software Wait for complete system recognition ⚠️ Critical Point: Do not proceed until all devices are recognized

### 3.2 Sample Loading

Objective Positioning Press “ESC” button on the front right side of the microscope body to lower objectives ⚠️ Safety: Ensure objectives are fully lowered before sample insertion

Plate Installation Remove plate holder piece Place sample plate on holder Insert holder assembly into microscope ⚠️ Critical Point: Check that the plate holder assembly is fully inserted on the stage and that the plate is not wobbly Press “ESC” button to raise objectives

Well Navigation Setup In Well plate and Slide Navigation panel: select correct well type from dropdown Double-click to navigate between wells or within wells ✅ Check: Z-position displayed at bottom right of screen

## 4 Phenotype Imaging Procedure

### 4.1 Initial Sample Localization

Time Required: 30 minutes Objective: 40X or 60X (water immersion)

Cell Finding Use Acquisition < Eyepieces < DAPI to locate cells ⚠️ Water Objective Setup: If using 40-60X water objective: Press “WID-Int” button to initiate water supply Remember to discharge water supply when done using “WID-Discharge”

Focus and Visualization Turn “PFS” (Perfect Focus System) ON Use Acquisition < Camera < DAPI < Live to focus Use “Keep Auto Scale” button for better visualization ✅ Quality Check: Cells clearly visible and in focus

### 4.2 Optical Configuration Setup

Custom Optical Configurations Click “Add New Optical Configuration” Set appropriate exposure time Configure excitation wavelength Set emission filter Save configuration for future use ⚠️ Critical Point: Right-click on configurations and save changes before acquisition

Z-Offset Calculation Focus on reference channel (typically one of the Illumina florophores for ISS or other fluorophores used for phenotype measurement) Record Z-position Switch to each additional channel Focus and record Z-position Calculate: Z-offset = (Z position of reference channel) - (Z position of other channel) 📝 Document: Record all Z-offsets for JOBS pipeline input

### 4.3 Phenotype Acquisition

JOBS Pipeline Selection In JOBS Explorer panel: select pre-set phenotype pipeline Example: “231226-Pheno-GFP-6well-CellVis-Alignment with WID 2.0” Configure acquisition parameters Input calculated Z-offsets

Mask Generation for ISS Alignment Change objective to 10X Navigate to center of well A1 Locate cells using DAPI channel Generate mask: Binary < Mask < Generate mask by threshold Save mask: Binary < Mask < Save Mask 📝 Document: Note mask location for ISS re-alignment

## 5 ISS Imaging Procedure

### 5.1 T7-ISS Setup

Time Required: 45 minutes Objective: 10X (or 4X for strong T7 signals)

Optical Configuration Selection Select pre-set T7-ISS configurations: MiSeq kit configurations NovaSeqX kit configurations Configure for each fluorophore being used

Signal Optimization Go to Acquisition < Desired Optical Configuration < Live Visualize cells in each fluorescent channel Signal Saturation Check: Pause live image for each channel Place cursor on signal area Check intensity value at bottom of screen ⚠️ Critical Point: Intensity must be <4095 to avoid saturation If saturated: adjust exposure time and/or lamp power Save configuration changes

Z-Offset Calculation for ISS Repeat Z-offset calculation process as in phenotype imaging Keep reference channel in perfect focus 📝 Document: Record ISS-specific Z-offsets

### 5.2 ISS JOBS Pipeline Configuration

Pipeline Selection Choose appropriate ISS pipeline: “OPS_PFS_v1” (PFS-based) “OPS_Software_AF_v1” (Software autofocus-based)

Capture Definition Setup Load optical configurations (DAPI and MiSeq/NovaSeqX) Right-click to select reference channel Input calculated Z-offsets for all channels

Plate Definition and Alignment Select plate type from database Re-align plate holder using center point option Load mask image from phenotype acquisition ✅ Quality Check: Verify alignment accuracy

Well and Point Selection Select wells for imaging Generate Points: Configure imaging area: Working area boundaries Border settings Point placement radius ⚠️ Critical Point: Use same area as phenotype imaging Confirm objective set to 10X

### 5.3 Acquisition Execution

Run ISS JOBS Execute imaging pipeline Monitor acquisition progress Repeat for each sequencing cycle 📝 Document: Log completion time for each cycle

## 6 System Shutdown Procedure

### 6.1 Post-Acquisition Cleanup

Water System Discharge (if used) Press “WID-Discharge” button Confirm complete water discharge

Sample Removal Press “ESC” to lower objectives Remove sample plate and holder Clean stage area if needed

### 6.2 System Power-Down

Software Shutdown Save all configurations and data Close NIS Elements software

Hardware Shutdown Turn OFF microscope and digital mirror switch Turn OFF main microscope switch Turn OFF lamp boxes and incubator Turn OFF main power switch

## 7 Quality Control and Validation

### 7.1 Critical Control Points

CCP 1: All system devices recognized during startup

CCP 2: PFS functioning properly for stable focus

CCP 3: Signal intensity <4095 in all channels

CCP 4: Successful mask generation and plate alignment

CCP 5: Consistent image quality across all acquisition points

### 7.2 Expected Results

Sharp, well-focused images across all channels

Proper exposure without saturation

Consistent illumination across field of view

Successful automated acquisition completion

## 8 Troubleshooting

| Problem | Possible Cause | Solution | Prevention |
| --- | --- | --- | --- |
| Poor focus/drift | PFS malfunction | Recalibrate PFS, check sample preparation | Regular PFS maintenance |
| Signal saturation | High exposure/lamp power | Reduce exposure time or lamp intensity | Pre-optimize settings |
| Uneven illumination | Misaligned light path | Check lamp alignment, clean optics | Regular optical alignment |
| Software crashes | Memory overload | Restart software, reduce acquisition area | Monitor system resources |
| Water objective issues | Air bubbles/dry objective | Re-initialize WID system | Proper WID maintenance |
| Poor signal | Photobleaching | Reduce exposure, use ND filters | Minimize light exposure |
| Plate misalignment | Incorrect mask/reference | Regenerate alignment mask | Verify mask quality |
| Autofocus failure | Poor contrast | Switch to manual focus or PFS | Optimize sample preparation |

## 9 Equipment Maintenance

### 9.1 Daily Maintenance

Clean stage and objectives with lens paper

Check water levels for immersion system

Verify lamp stability and alignment

Clean computer screen and workspace

### 9.2 Weekly Maintenance

Clean all optical surfaces

Check and clean filters

Verify PFS calibration

Update software if needed

### 9.3 Monthly Maintenance

Professional optical alignment check

Deep clean of stage and mechanics

Backup system configurations

Review and update optical configurations

## 10 References and Related Documents

### 10.1 Related SOPs

SOP_009a: Conventional Library Preparation

SOP_009b: T7 Library Preparation

SOP_010: SBS Illumina Sequencing

### 10.2 Equipment Manuals

Nikon Ti2 User Manual

NIS Elements AR Documentation

JOBS Explorer User Guide

## 11 Appendices

### 11.1 Optical Configuration Templates

Standard DAPI settings

Phenotype imaging parameters

T7-ISS kit-specific configurations

### 11.2 Z-Offset Calculation Worksheet

| Channel | Z-Position | Z-Offset from Reference |
| --- | --- | --- |
| DAPI (reference) | _____ μm | 0 |
| Channel 1 | _____ μm | _____ μm |
| Channel 2 | _____ μm | _____ μm |
| Channel 3 | _____ μm | _____ μm |